## Slide 1: POST MARKETING SURVEILLANCE OF HIV RAPID TESTS IN COTE D’IVOIRE
Dr CABLAN Arsher
MOH/NPHL/ PMS program coordinator
RTCQI Best practise meeting, Cape Town June 3-7, 2024

## Slide 2: OUTLINE
INTRODUCTION
OBJECTIVES
METHODOLOGY
RESULTS
WAY FORWARD

## Slide 3: INTRODUCTION
Post Marketing Surveillance (PMS)
Monitors the safety, quality and efficacy of an IVD after it has been released on the market
Protects human and public health by reducing risks through continuous monitoring of IVDs placed on the market.
Ensures that the manufacturer complies with its obligations to assess the risks and/or mitigate the risks.
PMS activity started in Cote d’Ivoire in December 2019 conducted by the NPHL

## Slide 4
INTRODUCTION
HTS sites algorithm
Confirmatory Laboratory algorithm

## Slide 5: OBJECTIVES
General : To verify the performance of HIV rapid tests before distribution and after distribution at the site of use.
Specific :
1- Evaluate any new batch coming in country
2- Carry out an inter-comparison of batches of HIV rapid tests from the national medical store
3- Evaluate the performance of HIV tests kits in districts and health facilities

## Slide 6
METHODOLOGY
RANDOM SAMPLING
At the central medical store: NPSP
Districts
Health centers (CHU, CHR, HG, ESPC)
- 25 per/year
- 100 tests/lot (Max 200 tests/ Site)
- Storage condition monitored (T°, Humidity rate): on site,
during transport after sampling, during storage until
evaluation at the NPHL
Central d’achat
2020, 2023
2019

## Slide 7
N° | Etablissement sanitaire | Determine | Bioline | Stat-Pak
1 | EPHR Abobo Nord | 1 | 0 | 0
2 | EPHR d’Abengourou | 1 | 0 | 0
3 | EPHR d’Odienné | 1 | 0 | 1
4 | EPHR de Bondoukou | 1 | 1 | 1
5 | EPHR de Bouaflé | 1 | 1 | 1
6 | EPHR de Daloa | 0 | 1 | 1
7 | EPHR de Dimbokro | 1 | 0 | 0
8 | EPHR de Divo | 1 | 1 | 1
9 | EPHR de Gagnoa | 1 | 1 | 1
10 | EPHR de Guiglo | 1 | 0 | 0
11 | EPHR de Katiola | 1 | 1 | 1
12 | EPHR de Korhogo | 1 | 1 | 1
13 | EPHR de Man | 1 | 1 | 1
14 | EPHR de San Pedro | 1 | 1 | 1
15 | EPHR de Séguela | 1 | 1 | 1
16 | EPHR de Touba | 1 | 1 | 1
17 | EPHR de Yamoussoukro | 1 | 1 | 1
18 | CHU de Bouake | 1 | 1 | 1
19 | CHU de Cocody | 1 | 1 | 1
20 | CHU de Treichville | 1 | 1 | 1
21 | CHU de Yopougon | 0 | 1 | 1
22 | DDSHP d’Abobo-Ouest | 1 | 0 | 1
23 | DDSHP de Dimbokro | 1 | 0 | 1
24 | DDSHP de Koumassi-Port-Bouët-Vridi | 1 | 2 | 1
25 | DDSHP de Treichville-Marcory | 1 | 1 | 1
26 | DDSHP de Yamoussoukro | 2 | 1 | 2
27 | ESPC Abidjan | 8 | 0 | 0
28 | Nouvelle PSP | 4 | 2 | 1
 | Total 16 EPHR, 4 DDSHP, 4 CHU, 8 ESPC | 37 | 22 | 24
COLLECTION SITE LOCATION

## Slide 8: METHODOLOGY
EVALUATION
Use of panels characterized by 2 ELISA techniques in parallel +/- LIA
(In house Panels, Commercial Panels)
Reference lot compared to subsidiary lots: Inter-lot comparison:
Panel maison à partir de poche de sang
In house panel from NBTS
Commercial panel
Characterisation algorithm

## Slide 9: METHODOLOGY
EVALUATION
 | Panel | Criteria | Accepted results
1 | Reference* (100 members) | Analytical Sensitivity anti-VIH-1 et anti-VIH-2 | ≥ 90%
2 | Dilution* (6 panels: 4 VIH1, VIH2, VIH1 and 2, 13 members each) | Analytical Sensitivity anti-VIH-1, anti VIH-2, anti VIH-1+VIH-2 Intra-analytical accuracy using the last positive | ≥ 90% ≥ 80%
3 | Seroconversion** (4 panels VIH1 de 21 membres) | Analytical Sensitivity from newly infected patients specimen | ≥ 90%
4 | Low title** (20 membres) | Analytical Sensitivity with sera with a low optical density/threshold value ratio | ≥ 80%
* In house Panel maison
** Commercial Panel

## Slide 10: METHODOLOGY
Performing rapid tests
Testing according to the manufacturer's recommendations
Reading of the results by two independent technicians
If discrepant results, then third reading by a supervisor, the final result is the one that is read twice out of the three.

## Slide 11: RESULTS

## Slide 12: INTERCOMPARISON OF BATCHES AT THE CENTRAL MEDICAL STORE
Chembio HIV 1/2 STAT-PAK could not be evaluated in a lot-to-lot comparison
because only one batch was available at the central warehouse.
However, this lot was used as a reference lot

## Slide 13: LOT COMPARISON : DETERMINE
The batches of Determine taken from the central warehouse comply with the acceptability criteria
defined in comparison to the reference lot
Limite d’acceptabilité
90%
Chart: 
| Analytical sensitivity using dilution panels of samples containing anti-HIV-1, anti-HIV-2 and anti-HIV-1+HIV-2 antibodies | Intra-analytical precision using the last positive of a panel of dilution series | Analytical sensitivity using samples with low anti-HIV-1 antibody titres | Analytical sensitivity using samples from HIV-1 seroconverting patients | Analytical sensitivity from undiluted reconverted sera containing anti-HIV-1, anti-HIV-2 antibodies
01191K200S | 0.96 | 1 | 0.95 | 1 | 1
01314K200R | 0.94 | 1 | 0.95 | 1 | 1
01321K200S | 1 | 1 | 0.9 | 1 | 1
90%
Determine reference batche : 00925K200R

## Slide 14: LOT COMPARISON : BIOLINE
The batche of Bioline taken from the central warehouse comply with the acceptability criteria
defined in comparison to the reference lot
Limite d’acceptabilité
Determine reference batche 03ADD009A
Chart: 
| Analytical sensitivity using dilution panels of samples containing anti-HIV-1, anti-HIV-2 and anti-HIV-1+HIV-2 antibodies | Intra-analytical precision using the last positive of a panel of dilution series | Analytical sensitivity using samples with low anti-HIV-1 antibody titres | Analytical sensitivity using samples from HIV-1 seroconverting patients | Analytical sensitivity from undiluted reconverted sera containing anti-HIV-1, anti-HIV-2 antibodies
 | 1 | 1 | 0.95 | 1 | 1
90%
80%

## Slide 15: POST-DISTRIBUTION RESULTS
With the low number of kits collected from the peripheral sites,
The comparison of batches used only dilution panel

## Slide 16: RESULTS
Chart: 
| January 2020 | October 2023 | Capital | Inland cities | University Hospitals | Regional Hospital | Health District | First contact-healthcare facility
Determine | 24 | 9 | 14 | 19 | 3 | 15 | 8 | 7
Bioline | 19 | 1 | 5 | 15 | 3 | 13 | 3 | 1
Stat-Pak | 23 | 0 | 7 | 16 | 5 | 13 | 5 | 0
PMS Sessions
Number of tests kits collected in health centers
Each year, 25 health centers are selected. At least 100 samples of each type of test
are expected (1 Determine kit, 4 Bioline kits, 5 Stat-Pak kits). On the 2020 and 2023 sessions, many
tests shortages have been observed in many health centers.

## Slide 17: RESULTS
Tests kits | Results (%) | Acceptance | Interpretation
Determine | 93.7-100 | ≥90% | Conforme
Bioline | 100
Stat-pak | 100
The tests collected and evaluated in post-distribution showed satisfactory performance according to the criteria defined using dilution panels

## Slide 18: RESULTS
Quality HIV kits batches assessed is acceptable both at the central medical stores and at the sites
No significant variation appeared between the batches evaluated.
Storage temperatures and humidity rate are within manufacturers' recommended ranges
Numerous tests kits stock-outs at the sites level

## Slide 19: WAY FORWARD
Development of a new database for the new HIV/Syphilis screening algorithms in Côte d'Ivoire
Characterization of an in-house panel in replacement of the commercial one expiring soon
Routine monitoring of any new test batch entering in the central medical store
Decentralization of some activities (PT, Certification, PMS, Misdiagnosis)
Develop the framework for PMS activity implementation including the report dissemination for decision making

## Slide 20: ACKNOWLEDGMENT

## Slide 21: THANKS YOU FOR ATTENTION